## tests
Execution | test name | PaceMaker_Mode | pchamber | polarity | Target | Unit | Min Limitation | Measurement | Max Limitation | Result | Min variation % | Variation | Max Variation % | Pluse Width(ms) | Pace_Voltage | Pace Rate | notice
Yes | pace gb voltage | DDD | V Path | bip | 1.5 | V | 1.35 |  | 1.65 |  | -10 |  | 10 | 0.5 | 1.5v | 40 | please_change_DDD_V_bip_1.5
Yes | pace gb pluse width | DDD | V Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 1.5v | 40 | please_change_DDD_V_bip_1.5
Yes | pace gb voltage | DDD | V Path | bip | 2 | V | 1.8 |  | 2.2 |  | -10 |  | 10 | 0.5 | 2v | 40 | please_change_DDD_V_bip_2.0
Yes | pace gb pluse width | DDD | V Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 2v | 40 | please_change_DDD_V_bip_2.0
Yes | pace gb voltage | DDD | V Path | bip | 2.5 | V | 2.25 |  | 2.75 |  | -10 |  | 10 | 0.5 | 2.5v | 40 | please_change_DDD_V_bip_2.5
Yes | pace gb pluse width | DDD | V Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 2.5v | 40 | please_change_DDD_V_bip_2.5
Yes | pace gb voltage | DDD | V Path | bip | 3 | V | 2.7 |  | 3.3 |  | -10 |  | 10 | 0.5 | 3v | 40 | please_change_DDD_V_bip_3.0
Yes | pace gb pluse width | DDD | V Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 3v | 40 | please_change_DDD_V_bip_3.0
Yes | pace gb voltage | DDD | V Path | bip | 3.5 | V | 3.15 |  | 3.85 |  | -10 |  | 10 | 0.5 | 3.5v | 40 | please_change_DDD_V_bip_3.5
Yes | pace gb pluse width | DDD | V Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 3.5v | 40 | please_change_DDD_V_bip_3.5
Yes | pace gb voltage | DDD | V Path | bip | 4 | V | 3.6 |  | 4.4 |  | -10 |  | 10 | 0.5 | 4v | 40 | please_change_DDD_V_bip_4
Yes | pace gb pluse width | DDD | V Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 4v | 40 | please_change_DDD_V_bip_4
Yes | pace gb voltage | DDD | V Path | bip | 5 | V | 4.5 |  | 5.5 |  | -10 |  | 10 | 0.5 | 5v | 40 | please_change_DDD_V_bip_5
Yes | pace gb pluse width | DDD | V Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 5v | 40 | please_change_DDD_V_bip_5
Yes | pace gb voltage | DDD | V Path | bip | 7.5 | V | 6.75 |  | 8.25 |  | -10 |  | 10 | 0.5 | 7.5v | 40 | please_change_DDD_V_bip_7.5
Yes | pace gb pluse width | DDD | V Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 7.5v | 40 | please_change_DDD_V_bip_7.5
Yes | pace gb voltage | DDD | V Path | unip | 1.5 | V | 1.35 |  | 1.65 |  | -10 |  | 10 | 0.5 | 1.5v | 40 | please_change_DDD_V_unip_1.5
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 1.5v | 40 | please_change_DDD_V_unip_1.5
Yes | pace gb voltage | DDD | V Path | unip | 2 | V | 1.8 |  | 2.2 |  | -10 |  | 10 | 0.5 | 2v | 40 | please_change_DDD_V_unip_2.0
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 2v | 40 | please_change_DDD_V_unip_2.0
Yes | pace gb voltage | DDD | V Path | unip | 2.5 | V | 2.25 |  | 2.75 |  | -10 |  | 10 | 0.5 | 2.5v | 40 | please_change_DDD_V_unip_2.5
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 2.5v | 40 | please_change_DDD_V_unip_2.5
Yes | pace gb voltage | DDD | V Path | unip | 3 | V | 2.7 |  | 3.3 |  | -10 |  | 10 | 0.5 | 3v | 40 | please_change_DDD_V_unip_3.0
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 3v | 40 | please_change_DDD_V_unip_3.0
Yes | pace gb voltage | DDD | V Path | unip | 3.5 | V | 3.15 |  | 3.85 |  | -10 |  | 10 | 0.5 | 3.5v | 40 | please_change_DDD_V_unip_3.5
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 3.5v | 40 | please_change_DDD_V_unip_3.5
Yes | pace gb voltage | DDD | V Path | unip | 4 | V | 3.6 |  | 4.4 |  | -10 |  | 10 | 0.5 | 4v | 40 | please_change_DDD_V_unip_4
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 4v | 40 | please_change_DDD_V_unip_4
Yes | pace gb voltage | DDD | V Path | unip | 5 | V | 4.5 |  | 5.5 |  | -10 |  | 10 | 0.5 | 5v | 40 | please_change_DDD_V_unip_5
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 5v | 40 | please_change_DDD_V_unip_5
Yes | pace gb voltage | DDD | V Path | unip | 7.5 | V | 6.75 |  | 8.25 |  | -10 |  | 10 | 0.5 | 7.5v | 40 | please_change_DDD_V_unip_7.5
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 7.5v | 40 | please_change_DDD_V_unip_7.5
Yes | pace gb pluse width | DDD | V Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 7.5v | 40 | please_change_DDD_A_unip_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.1 | 1.5v | 30 | please_change_DDD_V_30/mins bip_0.1ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.1 | 1.5v | 30 | please_change_DDD_V_30/mins bip_0.1ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.1 | 1.5v | 60 | please_change_DDD_V_60/mins bip_0.1ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.1 | 1.5v | 60 | please_change_DDD_V_60/mins bip_0.1ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.1 | 1.5v | 95 | please_change_DDD_V_95/mins bip_0.1ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.1 | 1.5v | 95 | please_change_DDD_V_95/mins bip_0.1ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.5 | 1.5v | 30 | please_change_DDD_V_30/mins bip_0.5ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.5 | 1.5v | 30 | please_change_DDD_V_30/mins bip_0.5ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.5 | 1.5v | 60 | please_change_DDD_V_60/mins bip_0.5ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.5 | 1.5v | 60 | please_change_DDD_V_60/mins bip_0.5ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.5 | 1.5v | 95 | please_change_DDD_V_95/mins bip_0.5ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.5 | 1.5v | 95 | please_change_DDD_V_95/mins bip_0.5ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 1 | 1.5v | 30 | please_change_DDD_V_30/mins bip_1ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 1 | 1.5v | 30 | please_change_DDD_V_30/mins bip_1ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 1 | 1.5v | 60 | please_change_DDD_V_60/mins bip_1ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 1 | 1.5v | 60 | please_change_DDD_V_60/mins bip_1ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 1 | 1.5v | 95 | please_change_DDD_V_95/mins bip_1ms_1.5
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 1 | 1.5v | 95 | please_change_DDD_V_95/mins bip_1ms_1.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.1 | 4v | 30 | please_change_DDD_V_30/mins bip_0.1ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.1 | 4v | 30 | please_change_DDD_V_30/mins bip_0.1ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.1 | 4v | 60 | please_change_DDD_V_60/mins bip_0.1ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.1 | 4v | 60 | please_change_DDD_V_60/mins bip_0.1ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.1 | 4v | 95 | please_change_DDD_V_95/mins bip_0.1ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.1 | 4v | 95 | please_change_DDD_V_95/mins bip_0.1ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.5 | 4v | 30 | please_change_DDD_V_30/mins bip_0.5ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.5 | 4v | 30 | please_change_DDD_V_30/mins bip_0.5ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.5 | 4v | 60 | please_change_DDD_V_60/mins bip_0.5ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.5 | 4v | 60 | please_change_DDD_V_60/mins bip_0.5ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.5 | 4v | 95 | please_change_DDD_V_95/mins bip_0.5ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.5 | 4v | 95 | please_change_DDD_V_95/mins bip_0.5ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 1 | 4v | 30 | please_change_DDD_V_30/mins bip_1ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 1 | 4v | 30 | please_change_DDD_V_30/mins bip_1ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 1 | 4v | 60 | please_change_DDD_V_60/mins bip_1ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 1 | 4v | 60 | please_change_DDD_V_60/mins bip_1ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 1 | 4v | 95 | please_change_DDD_V_95/mins bip_1ms_4
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 1 | 4v | 95 | please_change_DDD_V_95/mins bip_1ms_4
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.1 | 7.5v | 30 | please_change_DDD_V_30/mins bip_0.1ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.1 | 7.5v | 30 | please_change_DDD_V_30/mins bip_0.1ms_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.1 | 7.5v | 60 | please_change_DDD_V_60/mins bip_0.1ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.1 | 7.5v | 60 | please_change_DDD_V_60/mins bip_0.1ms_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.1 | 7.5v | 95 | please_change_DDD_V_95/mins bip_0.1ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.1 | 7.5v | 95 | please_change_DDD_V_95/mins bip_0.1ms_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.5 | 7.5v | 30 | please_change_DDD_V_30/mins bip_0.5ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.5 | 7.5v | 30 | please_change_DDD_V_30/mins bip_0.5ms_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.5 | 7.5v | 60 | please_change_DDD_V_60/mins bip_0.5ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.5 | 7.5v | 60 | please_change_DDD_V_60/mins bip_0.5ms_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.5 | 7.5v | 95 | please_change_DDD_V_95/mins bip_0.5ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.5 | 7.5v | 95 | please_change_DDD_V_95/mins bip_0.5ms_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 1 | 7.5v | 30 | please_change_DDD_V_30/mins bip_1ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 1 | 7.5v | 30 | please_change_DDD_V_30/mins bip_1ms_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 1 | 7.5v | 60 | please_change_DDD_V_60/mins bip_1ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 1 | 7.5v | 60 | please_change_DDD_V_60/mins bip_1ms_7.5
Yes | pace gb pluse frequency | DDD | V Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 1 | 7.5v | 95 | please_change_DDD_V_95/mins bip_1ms_7.5
Yes | pace gb pluse interval | DDD | V Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 1 | 7.5v | 95 | please_change_DDD_V_95/mins bip_1ms_7.5
Yes | pace gb voltage | DDD | A Path | bip | 1.5 | V | 1.35 |  | 1.65 |  | -10 |  | 10 | 0.5 | 1.5v | 40 | please_change_DDD_V_bip_1.5
Yes | pace gb pluse width | DDD | A Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 1.5v | 40 | please_change_DDD_V_bip_1.5
Yes | pace gb voltage | DDD | A Path | bip | 2 | V | 1.8 |  | 2.2 |  | -10 |  | 10 | 0.5 | 2v | 40 | please_change_DDD_V_bip_2.0
Yes | pace gb pluse width | DDD | A Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 2v | 40 | please_change_DDD_V_bip_2.0
Yes | pace gb voltage | DDD | A Path | bip | 2.5 | V | 2.25 |  | 2.75 |  | -10 |  | 10 | 0.5 | 2.5v | 40 | please_change_DDD_V_bip_2.5
Yes | pace gb pluse width | DDD | A Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 2.5v | 40 | please_change_DDD_V_bip_2.5
Yes | pace gb voltage | DDD | A Path | bip | 3 | V | 2.7 |  | 3.3 |  | -10 |  | 10 | 0.5 | 3v | 40 | please_change_DDD_V_bip_3.0
Yes | pace gb pluse width | DDD | A Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 3v | 40 | please_change_DDD_V_bip_3.0
Yes | pace gb voltage | DDD | A Path | bip | 3.5 | V | 3.15 |  | 3.85 |  | -10 |  | 10 | 0.5 | 3.5v | 40 | please_change_DDD_V_bip_3.5
Yes | pace gb pluse width | DDD | A Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 3.5v | 40 | please_change_DDD_V_bip_3.5
Yes | pace gb voltage | DDD | A Path | bip | 4 | V | 3.6 |  | 4.4 |  | -10 |  | 10 | 0.5 | 4v | 40 | please_change_DDD_V_bip_4
Yes | pace gb pluse width | DDD | A Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 4v | 40 | please_change_DDD_V_bip_4
Yes | pace gb voltage | DDD | A Path | bip | 5 | V | 4.5 |  | 5.5 |  | -10 |  | 10 | 0.5 | 5v | 40 | please_change_DDD_V_bip_5
Yes | pace gb pluse width | DDD | A Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 5v | 40 | please_change_DDD_V_bip_5
Yes | pace gb voltage | DDD | A Path | bip | 7.5 | V | 6.75 |  | 8.25 |  | -10 |  | 10 | 0.5 | 7.5v | 40 | please_change_DDD_V_bip_7.5
Yes | pace gb pluse width | DDD | A Path | bip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 7.5v | 40 | please_change_DDD_V_bip_7.5
Yes | pace gb voltage | DDD | A Path | unip | 1.5 | V | 1.35 |  | 1.65 |  | -10 |  | 10 | 0.5 | 1.5v | 40 | please_change_DDD_V_unip_1.5
Yes | pace gb pluse width | DDD | A Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 1.5v | 40 | please_change_DDD_V_unip_1.5
Yes | pace gb voltage | DDD | A Path | unip | 2 | V | 1.8 |  | 2.2 |  | -10 |  | 10 | 0.5 | 2v | 40 | please_change_DDD_V_unip_2.0
Yes | pace gb pluse width | DDD | A Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 2v | 40 | please_change_DDD_V_unip_2.0
Yes | pace gb voltage | DDD | A Path | unip | 2.5 | V | 2.25 |  | 2.75 |  | -10 |  | 10 | 0.5 | 2.5v | 40 | please_change_DDD_V_unip_2.5
Yes | pace gb pluse width | DDD | A Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 2.5v | 40 | please_change_DDD_V_unip_2.5
Yes | pace gb voltage | DDD | A Path | unip | 3 | V | 2.7 |  | 3.3 |  | -10 |  | 10 | 0.5 | 3v | 40 | please_change_DDD_V_unip_3.0
Yes | pace gb pluse width | DDD | A Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 3v | 40 | please_change_DDD_V_unip_3.0
Yes | pace gb voltage | DDD | A Path | unip | 3.5 | V | 3.15 |  | 3.85 |  | -10 |  | 10 | 0.5 | 3.5v | 40 | please_change_DDD_V_unip_3.5
Yes | pace gb pluse width | DDD | A Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 3.5v | 40 | please_change_DDD_V_unip_3.5
Yes | pace gb voltage | DDD | A Path | unip | 4 | V | 3.6 |  | 4.4 |  | -10 |  | 10 | 0.5 | 4v | 40 | please_change_DDD_V_unip_4
Yes | pace gb pluse width | DDD | A Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 4v | 40 | please_change_DDD_V_unip_4
Yes | pace gb voltage | DDD | A Path | unip | 5 | V | 4.5 |  | 5.5 |  | -10 |  | 10 | 0.5 | 5v | 40 | please_change_DDD_V_unip_5
Yes | pace gb pluse width | DDD | A Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 5v | 40 | please_change_DDD_V_unip_5
Yes | pace gb voltage | DDD | A Path | unip | 7.5 | V | 6.75 |  | 8.25 |  | -10 |  | 10 | 0.5 | 7.5v | 40 | please_change_DDD_V_unip_7.5
Yes | pace gb pluse width | DDD | A Path | unip | 0.5 | ms | 0.475 |  | 0.525 |  | -5 |  | 5 | 0.5 | 7.5v | 40 | please_change_DDD_V_unip_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.1 | 1.5v | 30 | please_change_DDD_V_30/mins bip_0.1ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.1 | 1.5v | 30 | please_change_DDD_V_30/mins bip_0.1ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.1 | 1.5v | 60 | please_change_DDD_V_60/mins bip_0.1ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.1 | 1.5v | 60 | please_change_DDD_V_60/mins bip_0.1ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.1 | 1.5v | 95 | please_change_DDD_V_95/mins bip_0.1ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.1 | 1.5v | 95 | please_change_DDD_V_95/mins bip_0.1ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.5 | 1.5v | 30 | please_change_DDD_V_30/mins bip_0.5ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.5 | 1.5v | 30 | please_change_DDD_V_30/mins bip_0.5ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.5 | 1.5v | 60 | please_change_DDD_V_60/mins bip_0.5ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.5 | 1.5v | 60 | please_change_DDD_V_60/mins bip_0.5ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.5 | 1.5v | 95 | please_change_DDD_V_95/mins bip_0.5ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.5 | 1.5v | 95 | please_change_DDD_V_95/mins bip_0.5ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 1 | 1.5v | 30 | please_change_DDD_V_30/mins bip_1ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 1 | 1.5v | 30 | please_change_DDD_V_30/mins bip_1ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 1 | 1.5v | 60 | please_change_DDD_V_60/mins bip_1ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 1 | 1.5v | 60 | please_change_DDD_V_60/mins bip_1ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 1 | 1.5v | 95 | please_change_DDD_V_95/mins bip_1ms_1.5
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 1 | 1.5v | 95 | please_change_DDD_V_95/mins bip_1ms_1.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.1 | 4v | 30 | please_change_DDD_V_30/mins bip_0.1ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.1 | 4v | 30 | please_change_DDD_V_30/mins bip_0.1ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.1 | 4v | 60 | please_change_DDD_V_60/mins bip_0.1ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.1 | 4v | 60 | please_change_DDD_V_60/mins bip_0.1ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.1 | 4v | 95 | please_change_DDD_V_95/mins bip_0.1ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.1 | 4v | 95 | please_change_DDD_V_95/mins bip_0.1ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.5 | 4v | 30 | please_change_DDD_V_30/mins bip_0.5ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.5 | 4v | 30 | please_change_DDD_V_30/mins bip_0.5ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.5 | 4v | 60 | please_change_DDD_V_60/mins bip_0.5ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.5 | 4v | 60 | please_change_DDD_V_60/mins bip_0.5ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.5 | 4v | 95 | please_change_DDD_V_95/mins bip_0.5ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.5 | 4v | 95 | please_change_DDD_V_95/mins bip_0.5ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 1 | 4v | 30 | please_change_DDD_V_30/mins bip_1ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 1 | 4v | 30 | please_change_DDD_V_30/mins bip_1ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 1 | 4v | 60 | please_change_DDD_V_60/mins bip_1ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 1 | 4v | 60 | please_change_DDD_V_60/mins bip_1ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 1 | 4v | 95 | please_change_DDD_V_95/mins bip_1ms_4
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 1 | 4v | 95 | please_change_DDD_V_95/mins bip_1ms_4
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.1 | 7.5v | 30 | please_change_DDD_V_30/mins bip_0.1ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.1 | 7.5v | 30 | please_change_DDD_V_30/mins bip_0.1ms_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.1 | 7.5v | 60 | please_change_DDD_V_60/mins bip_0.1ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.1 | 7.5v | 60 | please_change_DDD_V_60/mins bip_0.1ms_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.1 | 7.5v | 95 | please_change_DDD_V_95/mins bip_0.1ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.1 | 7.5v | 95 | please_change_DDD_V_95/mins bip_0.1ms_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 0.5 | 7.5v | 30 | please_change_DDD_V_30/mins bip_0.5ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 0.5 | 7.5v | 30 | please_change_DDD_V_30/mins bip_0.5ms_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 0.5 | 7.5v | 60 | please_change_DDD_V_60/mins bip_0.5ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 0.5 | 7.5v | 60 | please_change_DDD_V_60/mins bip_0.5ms_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 0.5 | 7.5v | 95 | please_change_DDD_V_95/mins bip_0.5ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 0.5 | 7.5v | 95 | please_change_DDD_V_95/mins bip_0.5ms_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 30 | min-1 | 29.85 |  | 30.15 |  | -0.5 |  | 0.5 | 1 | 7.5v | 30 | please_change_DDD_V_30/mins bip_1ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 2 | s | 1.996 |  | 2.004 |  | -0.2 |  | 0.2 | 1 | 7.5v | 30 | please_change_DDD_V_30/mins bip_1ms_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 60 | min-1 | 59.7 |  | 60.3 |  | -0.5 |  | 0.5 | 1 | 7.5v | 60 | please_change_DDD_V_60/mins bip_1ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 1 | s | 0.998 |  | 1.002 |  | -0.2 |  | 0.2 | 1 | 7.5v | 60 | please_change_DDD_V_60/mins bip_1ms_7.5
Yes | pace gb pluse frequency | DDD | A Path | bip | 95 | min-1 | 94.525 |  | 95.475 |  | -0.5 |  | 0.5 | 1 | 7.5v | 95 | please_change_DDD_V_95/mins bip_1ms_7.5
Yes | pace gb pluse interval | DDD | A Path | bip | 0.631 | s | 0.62974 |  | 0.63226 |  | -0.2 |  | 0.2 | 1 | 7.5v | 95 | please_change_DDD_V_95/mins bip_1ms_7.5